## Slide 1: Novel photon detectors
Peter Križan
University of Ljubljana and J. Stefan Institute

## Slide 2: Introduction: why?
Photon detectors are at the heart of most experiments in particle and nuclear physics. Moreover, they are also finding applications in other scientific fields (chemistry and biology) and are ubiquitous in society in general.
New environments where we need to detect light (in particular low light levels)  need advances in existing technology and transformative, novel ideas to meet the demanding requirements.
Two main lines of R&D to be pursued have been identified by the ECFA Detector R&D Roadmap:
- Enhance timing resolution and spectral range of photon detectors.
- Develop photosensors for extreme environments.
This talk: photosensors will be discussed in this spirit; also: the main emphasis will be on low light level detection.

## Slide 3: Contents
Introduction: why and what kind of photosensor?
Vacuum-based photodetectors
Solid state low light level photosensors
Gas-based photodetectors
Summary and outlook

## Slide 4: Vacuum-based photodetectors
Generic steps:
Photon  photoelectron conversion
Photoelectron collection in the multiplication system
Mutiplication (dynodes, microchannel plates, high E field + Si)
Signal collection

## Slide 5: Multianode photomultiplier tubes (PMTs)
Pioneered in the HERA-B RICH, later used in
RICH detectors of COMPASS, CLASS12 and GlueX
Recent use in the upgraded LHCb RICH detectors;
planned for CBM RICH
Excellent performance (excellent single photon detection efficiency, very low noise, low cross-talk), best choice for large areas with no B field
Metal channel dynodes
Response to single photons

## Slide 6: Micro Channel Plate PMT (MCP-PMT)
Multiplication step: a continuous dynode – a micro-channel

## Slide 7: Micro Channel Plate PMT (MCP-PMT)
MCP is a thin glass plate with an array of holes (<10-100 mm diameter) - - a continuous dynode structure
MCP gain depends on L/D ratio – typically 1000 for L/D=40
~ 400 mm
F ~ 10 mm
Similar to ordinary PMT – the dynode structure is replaced by MCPs.
Basic characteristics:
Gain ~ 106 → single photon
Collection efficiency ~ 60%
Small thickness, high field → small TTS
Works in magnetic field
Segmented anode → position sensitive
PHOTONIS
HAMAMATSU
Anodes → can be segmented according to application needs
window with photocathode
dual MCP -
Chevron config.
PHOTEK
LAPPD

## Slide 8: MCP-PMT: single photon pulse height and timing
Gain in a single channel saturates at high gains due to space charge effect → peaking distribution for single photoelectrons
Typical single photon timing distribution with a narrow main peak (s ~ 40 ps) and contributions from photoelectron elastic (flat distribution) and inelastic back-scattering.
Photoelectron back-scattering produces a rather long tail in timing distribution and position resolution.
Photoelectron backscattering reduces collection efficiency and gain, and contributes to cross-talk in multi-anode PMTs
S.Korpar, talk at PD07

## Slide 9
Photoelectrons travel from photocathode to the electron multiplier (uniform electric field , initial energy ):
photoelectron range
and maximal travel time (sideway start)
time difference between downward and sideways initial direction
Example (, , )
photoelectron:
max range
p.e. transit time
backscattering:
max range
max delay
Backscattering delay and range (maximum for elastic scattering):
maximum range vs. angle maximum range for backscattered photoelectron is twice the photocathode – first electrode distance
maximum delay vs. angle maximum delay is twice the photoelectron travel time
time of arrival of elastically scattered photoelectrons: flat distribution up to max
Note: the maximum range is dramatically reduced in high B-field perpendicular to the MCP-PMT window
Photoelectron in uniform electric field
S. Korpar, Talk at PD07
MCP1
window with photocathode

## Slide 10: MCP PMT modeling
Tails understood (elastic and inelastic scattering of photoelectrons off the MCP), can be significantly reduced by:
decreased photocathode-MCP distance and
increased photocathode-MCP voltage difference
Timing: Expect to improve with increased photocathode-MCP voltage difference as sqrt(HVPC-MCP1)
MCP-PMT LAPPD – timing vs PC-MCP1 voltage 
s(ps)

## Slide 11: LAPPD (large area picosecond photodetector) Gen II
Characteristics (Incom):
borosilicate back plate with interior resistive ground plane anode – 5 mm thick
capacitively coupled readout electrode
MCPs with 20 m pores at 20 m pitch
two parallel spacers (active fraction %)
gain @ ROP (825 V/MCP, 100 V on photocathode)
peak QE
size 230 mm x 220 mm x 22 mm (243 mm X 274 mm X 25.2 mm with mounting case)
Dark Count rate @ ROP: ~ 70 kHz/cm2 with 8x105 gain

## Slide 12: LAPPD – timing distribution
measured timing distribution typical for MCP-PMT
main prompt peak with some inelastic and elastic backscattering contribution
S. Korpar et al., to be submitted to NIMA
s(ps)
LAPD timing: sigma of the main peak 27 ps
LAPPD – timing vs PC-MCP1 voltage

## Slide 13
MCP PMT readout: capacitive coupling vs. internal anodes
LAPPD (capacitive coupling through the backplane)
mm
mm
BURLE/Photonis PLANACON (internal anodes)
mm
mm
mm
mm
mm
Capacitive coupling vs. internal anodes: signal spread comparison for two MCP PMTs with the same pad size, same range: charge sharing is more effective for capacitive coupling (spreads over larger area).
Advantage or not? Depends on the usage
Secondary electrons spread out when traveling from the MCP-out electrode to the anode and can hit more than one anode → Charge sharing
Fraction of charge detected by the right pad as a function of red laser spot position

## Slide 14
MCP PMT readout: capacitive coupling, modeling
LAPPD charge sharing
S. Korpar, DRD4 Coll. Meeting April 2025, to be published

## Slide 15: MCP PMT ageing
MCP PMTs: photo-cathode degradation due to ion feedback, main concern in high intensity experiments
ALD (atomic layer deposition) coating of MCP pores → ~100x photo-cathode lifetime increase
Hamamatsu 1-inch YH0205 (>20 C/cm²) [K. Inami, 2021]
No QE degradation for Photonis MCP-PMT (R2D2) to >34 C/cm²
Little QE degradation in LAPPD 8-inch up to 5.6 C/cm² [V. A. Chirayath, CPAD2021]
LAPPD-64
K. Inami, 2021, Talk at ECFA TF4 Symposium
A. Lehman et al.,
https://arxiv.org/abs/2403.13938
V.A. Chirayath et al., Talk at CPAD 2021

## Slide 16: MCP PMTs: high rate operation
  High rate performance studies
A. Lehman et al.,
https://arxiv.org/abs/2403.169536

## Slide 17
4.2.1: VPD: New material, new coatings, longevity and rate capability study
  This concerns the R&D on new materials to produce VPD,
  new shapes and new coatings and their consequences
  on their longevity and rate capability
4.2.2: VPD-PMT: New photocathode materials, structure and high quantum efficiency VPD
  New photocathode materials, new structures and their impact on improving
  the quantum efficiency for different wavelengths
4.2.3: VPD time and spatial resolution performance
  Study of VPD timing and spatial performance using appropriate readout electronics and appropriate anode structures
Amorphous Si MCPC(Geneva)
[1] A. Franco et al. Sci. Rep. 1 (2014).
Si nanometric structure for
reflective photocathode (Lyon)
MCP+Timepix4 (Ferrara)
MCP+PICMIC concept (Lyon)
time
Woven strips
MCP-PMTs: New materials and read-out options

## Slide 18: Possible Future of Electron Multiplication
Tynodes (→ Time Photon Counter)
Transmission mode dynode → tynode
Fabrication of tynodes (MgO ALD, diamond) using MEMS technology
“Anode” is a CMOS chip (e.g., TimePix) Very promising properties
Very compact; high B-field tolerance; very fast
Very low DCR; very good 2D spatial resolution
MCP-PMT with CMOS anode
Conceptual design for 4D detection of single photons
Hybrid concept: MCP-PMT where the pixelated anode is an
ASIC (CMOS) embedded inside the vacuum
Prototype with Timepix4 ASIC as anode (array of 23k pixels)
Envisaged performance
<100 ps time resolution and 5-10 μm spatial resolution
Rate capability of >100 MHz/cm² (<2.5 Ghits/s @ 7 cm² area)
Low gain (~104) operation possible → x100 lifetime increase
M. Fiorini, RICH2022
H. van der Graaf et al., NIM A847 (2017) 148

## Slide 19: Hybrid photodetectors (HPD, HAPD)
Photo-electron acceleration in a static electric field (8kV to 25 kV)
Photo-electron detection with
Segmented PIN diode (HPD)
Avalanche photo diode (HAPD)
Silicon photomultiplier (VSiPMT)
Employed on a large scale:
HPD: RICH1+RICH2 of LHCb (Run 1+2), CMS HCAL
HAPD: Aerogel RICH detector of Belle II
Focusing and proximity focusing configurations

## Slide 20
around 20% of photoelectrons back-scatter and the maximum range is twice the distance from photocathode to APD ~40mm
in a strong magnetic field (perp. to the HAPD window) scattered photoelectrons follow magnetic field lines and fall back to the same pad
photoelectron energy is deposited at the same pad
-- 0 T
-- 1.5 T
0 T
1.5 T
HAPD: photoelectron backscattering in magnetic field

## Slide 21
An array of APDs operated in Geiger mode – above APD breakdown voltage (microcells or SPADs – single photon avalanche diodes)
Detection of photons:
absorbed photon generates an electron-hole pair
an avalanche is triggered by the carrier in the high field region → signal
voltage drops below breakdown and avalanche is quenched (passive or active quenching)
each triggered microcell contributes the same amount of charge to the signal
Solid state low light level photosensors:
Silicon photomultipliers SiPM

## Slide 22: SiPMs as single photon detectors
SiPM as low-light-level sensors.
Characteristics:
low operation voltage ~ 10-100 V
gain ~ 106
peak PDE up to 65%(@400nm) PDE = QE x εgeiger x εgeo (up to 5x PMT!)
εgeo – dead space between the cells
time resolution ~ 100 ps
works in high magnetic field
dark counts ~ few 100 kHz/mm2
radiation damage (p,n)
*PDE = photon detection efficiency
  050U
  100U
  025U
few mm

## Slide 23
SiPM: noise
A. Gola et al. Sensors 19(2019)308
dark counts are produced by thermal generation of carriers, trap assisted tunnelling or band gap tunnelling
signal equal to single photon response
typical rate went from for early attempts to below for more recent devices
roughly halved for every -
increases linearly with fluence
optical cross-talk produced when photons emitted in avalanche initiate signal in neighbouring cell, reduced by screening – trenches
after-pulses produced by trap-release of carriers or delayed arrival of optically induced carrier in the same cell

## Slide 24: SiPM: parameter correlation
Higher overvoltage:
higher field:
  higher avalanche trigger probability → higher PDE
  faster signal → better timing
higher gain:
  better signal to noise (electronic)
  more optical cross-talk → higher ENF, worse timing
  more after-pulses
A. Gola et al. Sensors 19(2019)308

## Slide 25
VUV SiPM for cryogenic applications
A. Gola et al. Sensors 19(2019)308
A. Jamil et al. IEEE TNS 65-11(2018)2823
LAr, LXe applications:
VUV sensitivity required:
  128 nm (LAr), 178 nm (LXe)
  optimization of anti-reflective coating ARC
  PDE
cryogenic temperatures:
  low DCR dominated by band-band tunnelling, reduced by low-field avalanche region
  higher after-pulse rate

## Slide 26
SiPM: single photon timing
F. Acerbi et al. IEEE TNS 61(2014)2678
Intrinsic TTS of SiPM microcells is extremely fast, for single microcells (SPAD), but timing deteriorates for larger devices. The main contributions:
nonuniformity within microcell (edges)
spread between microcells
overall SiPM capacitance
dependence - tails
Comparison of timing properties for single SPAD, and SiPMs with the same SPAD for microcells:
  timing improves with higher overvoltage – larger pulses, at the expense of increased SiPM noise
  best timing resolutions for single cell signals are
  TTS deterioration mainly due to a larger overall capacitance  reduced signal slope,

## Slide 27: SiPM: timing variation
KETEK PM3375TS-SBO (early design)
Variation of TTS over the device surface can contribute to overall time spread:
  variation within micro-cell
  variation for different micro-cells
.
S. Korpar et al. @IEEE 2015
F. Acerbi et al. NIM A926(2019)16
0.3 mm
2.5 mm
FBK: Masking of outer regions of micro-cells: Improve signal peaking and mask areas of micro-cell with worse timing

## Slide 28
SiPM: timing for multi-cell signals
F. Acerbi et al. IEEE TNS 61(2014)2678
Optical cross-talk contribution to multi-cell signals spoils timing distribution – does not scale with :
two components for 2-micro-cell signals:
  double photon events – proper scaling
  single photon with cross-talk, timing somewhere between single and double micro-cell signals and resolution is worse
ratio between contributions changes with light intensity confirming optical cross-talk origin
even more components for multi-micro-cell signals

## Slide 29
SiPM: timing test with pico-second laser
AdvanSiD SiPM ASD-NUV3S-P-40
OV=6V, T=-25o C
blue laser l =408nm, ~ 35ps FWHM
R. Dolenec et al. NIM A876(2017)257

## Slide 30: Reduction of optical crosstalk
~ few µm
High-field region
Quenching resistor
Metal
Connections
Starting from the NUV-HD technology, FBK and Broadcom jointly developed the NUV-HD-MT technology, adding metal-filled deep trench isolation to strongly suppress optical crosstalk.
Other changes: low electric field variant, layout optimized for timing.
Microcell (SPAD) of SiPM
Metal-filled DTI
Conceptual drawing of the NUV-HD-MT, with the addition of metal-filled Deep Trench Isolation.
Reduction of optical crosstalk probability in NUV-HD-MT, compared to the “standard” NUV-HD. Measurement without encapsulation resin, i.e. only considering internal crosstalk probability.
> 10x
reduction
A. Gola, talk at RICH2022

## Slide 31: Backside illuminated SiPMs
The first results of the FBK IBIS Run samples
Backside illuminated (BSI) SiPMs:
potential for an enhanced PDE and a better radiation tolerance.

## Slide 32: Ultragranular SiPMs with integrated electronics
256 Ch 5x5cm2
Module proposal
FastIC+
(collab. CERN ICCUB)
Low power
Chip
Timing of SSPD & Developing ultra-granular SiPM that integrates with the readout electronics
design & production 6.2025
2024 Produced and evaluated
3x3 mm2 SiPM array
Long term goal 1x1mm2 SiPM array
Ultra granular SiPMs

## Slide 33: SiMs: SPAD with micro-lenses
CMOS-SPAD light sensors: co-integration of SPADs and electronics, digitised output signals
spadRICH - Radiation-hard digital analog silicon photomultipliers for future upgrades of Ring Imaging Cherenkov detectors

## Slide 34: Light concentrators
At the device level (lenses, Winston cones):
reduce active are – reduce DCR (tolerate higher fluences)
use smaller faster devices
At the micro-cell level (micro-lenses, diffractive lenses, meta lenses):
compensate for low fill factor – small cells, dSiPM
concentrate light in cell centre – better timing
Higher concentration – narrower angular acceptance
Imaging light concentrators:
smaller photon impact angles on the sensor
can be used with position-sensitive arrays
Non-imaging light concentrators:
larger photon impact angles on the sensor – directly coupled to the sensor

## Slide 35
SiPM RICH with light concentrators
RICH photon detector module prototype:
Hamamatsu 64 channel MPPC module S11834-3388DF, array of SiPMs @ pitch
matching array of quartz light concentrators used
two thick aerogel tiles in focusing configuration ()
tested in electron beam at DESY
E. Tahirović et al., NIM A787 (2015) 203

## Slide 36: Microlenses for light collection
J.M. Pavia et al. Opt.Exp. 22-4(2014)4202
Micro-lens array coupled to SPAD array
CMOS SPAD array, 128x128 diameter @ pitch – 5% fill factor
matching polymer plano-convex micro-lens array

## Slide 37: Microlenses for light collection 2
Increase the active area fraction – use microlenses
G. Haefeli et al., TNS, DOI 10.1109/TNS.2025.3542597

## Slide 38: SiPMs: Radiation damage
Show stopper at fluences above ~1011 in case single (or few) photon sensitivity is required!
(e.g. expected fluence in the ARICH area of Belle II: 2-20 1011 n cm-2 )
Use of wave-form sampling readout electronics
Operating the SiPMs at lower temperature
Annealing periodically (annealing at elevated temperature is preferred)
Reducing recovery time to lower cell occupancy
 Radiation resistant SiPMs, other materials?
Waveforms from irradiated SiPMs

## Slide 39
Beyond 107÷108 neq/cm2 little correlation between the DCR before and after irradiation:
All technologies seem to “converge” towards similar values
Independence of bulk damage from contaminants in the SiPM starting material?
VOV = 3 V
SiPMs: Radiation damage
A. Gola, RICH2022
Room temperature annealing (20-25°C) for samples irradiated to 6.4∙1011 1 MeV neq/cm2
Little effect, knee point at around 1.5·103 min (~1 day)
A. Gola, RICH2022
DCR (dark count rate) vs fluence

## Slide 40: SiPMs: Radiation damage, annealing at elevated temperatures
M. Calvi et al., NIMA 922 (2019) 243-249
Dark counts at -30C of a Hamamatsu S13360-1350CS SiPMs: non irradiated (blue) and irradiated up to 1011 (yellow), 1012 (green) and 1013 (orange) neq/cm2
annealing

## Slide 41: SiPMs: radiation damage, mitigation
Operation of highly irradiated SiPMs for single photon detection – cooling and annealing
Cooling: Temperature at which the SiPMs are ”usable” for single-photon detection, i.e. where the single photo electron peak @ 9V over-voltage is separated from the background.
LHCb RICH Upg II -2034
Belle II ARICH
Upg ~2032
EIC dRICH
Upg ~2031
Annealing
R. Preghenella, PD24
D. Consuegra-Rodrigez et al., Eur. Phys. J. C 84, 970 (2024)

## Slide 42: SiPMs: Radiation damage, annealing at elevated temperatures
 Annealing helps also at 77 K
M. Calvi et al., NIMA 922 (2019) 243-249
DCR at 77 K versus neutron fluence
Blue circles: annealed sample

## Slide 43: SiPMs: radiation damage, mitigation
Fast & radiation hard SiPMs - enabling the use of SiPM in highly irradiated areas
Experimental structures, AidaInnova Run – exp. May 2025
Two different technologies:
  Low electric field
  Ultra Low electric field
Cell pitch: 15, 25, 40, 75um; SiPM sizes: (0.25, 0.5,1,2,3)2 mm2

## Slide 44: New materials
new higher band gap materials - possibly lower DCR - higher radiation resistance, higher temperature
(V)UV sensitive
high dark count rate – dominated by trap assisted tunnelling
L. Su et al. J. Semicond. 40(2019)121802
4H-SiC:
PDE
DCR
nonuniform response

## Slide 45: Gas based photon detectors
Standard photosensitive substance: CsI evaporated on one of the cathodes.
Large scale application: ~11 m2 ALICE RICH

## Slide 46
Gas based photon detectors: recent developments
Instead of MWPC:
Use chambers with multiple GEM or thick GEM (THGEM) gas amplification stages with transm. or refl. photocathode
COMPASS RICH: trans. photocathode and 2x THGEM + MicroMegas
Ion damage of the photocathode:
blocking ions – non-aligned GEM holes
New developments:
Smaller pads
Novel photocathode material: nano-diamond layer
S. Dalla Torre, NIM A 970 (2020) 163768

## Slide 47: Summary and outlook
Next generation of experiments in particle physics: photosensors with faster timing, wider spectral range and improved radiation tolerance.
Many new interesting developments are underway, in particular in SiPMs and MCP-PMTs – not all of them could be covered in this talk.
A detector R&D collaboration (DRD4) was set up early in 2024 to facilitate collaboration in this area of research

## Slide 48: Back-up slides

## Slide 49: ECFA Detector R&D Roadmap
The ECFA Detector R&D Roadmap, developed following the 2020 European Strategy for Particle Physics, outlines a long-term vision to advance detector technologies critical for future particle physics experiments.
It emphasizes strategic planning and investment in areas like
sensor development (gaseous, liquid, solid-state),
photon detection,
quantum sensing,
calorimetry, and
integrated electronics.
The roadmap highlights the need for coordinated European efforts, robust infrastructure, training, and industrial partnerships.
Strategic recommendations address challenges such as rising R&D costs, sustainability, and the retention of expert talent to ensure Europe remains a global leader in detector innovation.
It also provides a list of detector research and development themes - DRDTs

## Slide 50: Detector research and development themes (DRDTs)

## Slide 51: ECFA Detector R&D Roadmap implementation: Detector R&D (DRD) Collaborations
1. Gaseous
2. Liquid
5. Quantum
6. Calorimetry
7. Electronics
8. Integration
e.g. time/spatial resolution;
environment friendly gases
e.g. Light/charge readout;
low background materials
3. Semiconductor
e.g.
CMOS pixel sensors;
High time resolution
(10s ps)
4. PID & Photon
e.g.
spectral range
of photon sensors;
Time resolution
quantum sensors
- R&D, incl. beyond QFTP in conventional
detectors
e.g. Sandwich; noble liquid; optical
e.g. ASICs; FPGAs; DAQ
tracking detector mechanics
Chris Parkes, IOP APP+HEPP+NP Conference, April 2024

## Slide 52: DRD4: photon detectors and PID
Detector R&D Themes (DRDTs)
  Single-photon sensitive photodetectors (vacuum, solid state, hybrid)
  PID techniques (Cherenkov-based, Time of Flight)
  Scintillating Fiber (SciFi) tracking
  Transition Radiation (TR) using solid state X-ray detectors

## Slide 53: DRD4 : photon detectors and PID
Organization:
Work packages: projects
Working groups: discussion forums

## Slide 54: DRD4:WP4.1 Solid-State Photodetectors
◼ Task 1 -SSPD with new configurations and modes: Development of back-side illuminated SiPM (potential for better PDE and radiation tolerance); development of ultra-granular SiPM that integrates with the electronics by using 2.5D or 3D interconnection techniques; development of CMOS-SPAD light monolithic sensors for HEP; study of new materials for light detection
◼ Task 2 -Fast radiation hard SiPMs: Standardize procedures for quantification of radiation effects; irradiated SiPMs characterization in wide temperatures range (down to -200 ˚C); study of annealing; study and quantify other measures enabling the use of SiPM in highly irradiated areas (e.g. smaller SiPMs, macro-and micro-light collectors)
◼ Task 3 -Timing of SSPD, including readout electronics: Study and improve the timing of SiPMs; co-design of a multi-ch. readout ASIC exploiting the timing potential; integration and packaging with integrated cooling; vertical integration of SiPM arrays to FEE (better timing via reduction of interconnections' parasitic induct+capac)
Long term goal 1x1mm2 SiPM array

## Slide 55: DRD4:WP4.2 Vacuum-based Photodetectors
◼Task 1 -New materials, coatings, longevity and rate capability studies: Develop new materials and techniques to increase MCP-PMT tube lifetime and improve rate capabilities; use new techniques with new materials to achieve high aspect ratio with small diameter for better gain, time, and spatial resolution
◼Task 2 -New photocathode materials, structure and high QE VPD: Search for new materials with the required characteristics to be used as photocathodes; develop photocathodes with new structures
◼Task 3 -VPD time and spatial resolution performance: Development of large area MCP-based photodetector with combined excellent timing and position resolution, including electronics integration

## Slide 56: Rate Capability
MaPMT
MaPMT R13742: no gain loss to up 10 MHz/mm²
MCP-PMTs
Gain decreases at high photon rates
2-inch MCP-PMTs: ~1 MHz/cm² (with 106 gain)
1-inch Hamamatsu R10754 PMTs: ≥10 MHz/cm²
8-inch LAPPD-64: ~100 kHz at 4.6 mm Ø laser spot
Improving the rate capability in MCP-PMTs
Lower MCP resistance (ranging at tens of MΩ)
Lower capacitance - subdivision of MCP layers into smaller partial areas; the second MCP layer was divided in the previous version of Hamamatsu tubes
Low gain operation (for single photon detection) - may be possible with amplifiers included directly in the anode (e.g., monolithic or CMOS designs)
R10754 MCP-PMT
LAPPD-64
A. Lehmann, RICH2022
M. Calvi et al., JINST 15 (2020) P10031
V.A. Chirayath et al., Talk at CPAD 2021

## Slide 57: Radiation Tolerance
MaPMT
Recently tested for CBM (H12700 and H8500) Irradiation dose up to 3x1011 n/cm² and 150 Gy gammas Few kHz increase of DCR but recovers after few months
UV glass: 1-2% transmission loss @ 400 nm up to 1 kGy No negative effects on gain and pulse height spectrum Neutron activation of Kovar metal alloy is a minor effect
H(A)PD
Radiation tolerance dominated by silicon detector Intensively studied for 144ch HAPD for ARICH @Belle II
Basically no QE loss at 5x1011 n/cm²
Increase of APD leakage current leads to significantly worse S/N ratio: ~17 before, ~3 after irradiation
S/N ratio partly regained by adjusted readout parameters
MCP-PMT
Entrance window: should be similar to other PMTs
MCPs: currently no data
Pulse height dist.
after 5x1011 n/cm²
Pulse height dist.
after 5x1011 n/cm²
+ adjusted readout
C. Pauly et al., NIM A1040 (2022) 167177
.S. Shiizuka, Talk at VCI 2010
I. Adachi, NIM A639 (2011) 103

## Slide 58
SiPM: p-on-n vs. n-on-p
n-on-p: green/red light sensitive:
electrons drift to Geiger region from substrate and holes from surface side
higher dark count rate – most of the thermally generated carriers arriving to Geiger region are electrons
p on n - green/blue light sensitive:
electrons drift to Geiger region from surface and holes from substrate side
lover dark count rate – most of the thermally generated carriers arriving to the Geiger region are holes
J. Musienko@PD07
holes
electrons
holes
electrons
A. Gola et al. Sensors 19(2019)308